## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S22475231
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 174.0791 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.14465 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 693.1794 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1733 | 0.4987 | 0.03 | 0.018083 | PASS
Radius Red | 0.1733 | 0.5284 | 0.03 | 0.00080623 | PASS
Radius Green | 0.1733 | 0.5621 | 0.03 | 0.011 | PASS
Radius Blue | 0.1733 | 0.3047 | 0.03 | 0.022066 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.07397 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 58.5478 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 64.3171 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1733 | 0.4987 | 0.018083 | PASS | -0.0037 | 0.0177 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4517 | 0.5284 | 0.00080623 | PASS | 0.0007 | 0.0004 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.123 | 0.5621 | 0.011 | PASS | 0.011 | 0.0001 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.119 | 0.3047 | 0.022066 | PASS | -0.004 | 0.0217 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 235.7965
Point 1 | 210.3752
Point 2 | 203.8958
Point 3 | 213.0312
Point 4 | 210.7255

## Curves
CURVES
Step | Luminance | Status
1 | 0.11079 | 0
2 | 0.050434 | 0
3 | 0.080671 | 0
4 | 0.065553 | 0
5 | 0.073141 | 0
6 | 0.073141 | 0
7 | 0.10323 | 0
8 | 0.15448 | 0
9 | 0.072149 | 0
10 | 0.11359 | 0
11 | 0.13443 | 0
12 | 0.1241 | 0
13 | 0.12416 | 0
14 | 0.16286 | 0
15 | 0.16289 | 0
16 | 0.21312 | 0
17 | 0.21321 | 0
18 | 0.27342 | 0
19 | 0.27345 | 0
20 | 0.35315 | 0
21 | 0.4089 | 0
22 | 0.19864 | 0
23 | 0.30383 | 0
24 | 0.35695 | 0
25 | 0.3301 | 0
26 | 0.3301 | 0
27 | 0.42554 | 0
28 | 0.42525 | 0
29 | 0.54841 | 0
30 | 0.54812 | 0
31 | 0.70135 | 0
32 | 0.70018 | 0
33 | 0.8931 | 0
34 | 0.89339 | 0
35 | 1.1494 | 0
36 | 1.1482 | 0
37 | 1.4605 | 0
38 | 1.4611 | 0
39 | 1.8656 | 0
40 | 1.8662 | 0
41 | 2.3702 | 0
42 | 2.3708 | 0
43 | 3.012 | 0
44 | 3.0441 | 0
45 | 3.8643 | 0
46 | 3.8497 | 0
47 | 4.9062 | 0
48 | 4.9091 | 0
49 | 6.2634 | 0
50 | 6.2138 | 0
51 | 8.0175 | 0
52 | 8.0905 | 0
53 | 10.2824 | 0
54 | 11.8467 | 0
55 | 5.9219 | 0
56 | 8.7559 | 0
57 | 10.2969 | 0
58 | 9.541 | 0
59 | 9.5293 | 0
60 | 12.162 | 0
61 | 12.0043 | 0
62 | 15.5855 | 0
63 | 15.1798 | 0
64 | 19.298 | 0
65 | 18.977 | 0
66 | 23.7373 | 0
67 | 29.8285 | 0
68 | 15.9095 | 0
69 | 23.1448 | 0
70 | 27.7825 | 0
71 | 24.7996 | 0
72 | 25.7511 | 0
73 | 33.0098 | 0
74 | 31.317 | 0
75 | 42.4078 | 0
76 | 38.526 | 0
77 | 51.8058 | 0
78 | 51.6015 | 0
79 | 65.8736 | 0
80 | 66.1946 | 0
81 | 83.9983 | 0
82 | 84.1734 | 0
83 | 106.0632 | 0
84 | 106.4426 | 0
85 | 135.9792 | 0
86 | 134.4615 | 0
87 | 213.469 | 0
88 | 106.1216 | 0
89 | 159.4742 | 0
90 | 186.1797 | 0
91 | 172.6081 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 191.0539
150 | 200 | 174.0791 | PASS | Point 1 | 170.3024
|  |  |  | Point 2 | 165.2823
|  |  |  | Point 3 | 172.6665
Uniformity Test |  |  |  | Point 4 | 171.0904
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.14465 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 191.0247 | White
100 | 693.1794 | PASS | 0.27558 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.081313
0.05 | 0.1 | 0.07397 | PASS | Point 1 | 0.072382
|  |  |  | Point 2 | 0.070281
|  |  |  | Point 3 | 0.073462
|  |  |  | Point 4 | 0.072411

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 191.1122 | 73.3453 | 172.0827 | 64.3559
50 | 58.5478 | PASS | Point 1 | 170.5358 | 67.9166 | 165.8368 | 66.1363
|  |  | Point 2 | 165.4282 | 66.4865 | 162.1594 | 64.064
|  |  | Point 3 | 172.5789 | 74.863 | 159.124 | 60.1823
|  |  | Point 4 | 171.0904 | 72.4697 | 159.7953 | 58.5478

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 190.8787 | 73.4037 | 171.9368 | 64.4143
20 | 64.3171 | PASS | Point 0 | Black | 0.27704 | 0.34965 | 0.29102 | 0.98913
|  |  | Point 1 | White | 170.4483 | 68.0042 | 165.866 | 66.253
|  |  | Point 1 | Black | 0.2494 | 0.32105 | 0.28066 | 1.0063
|  |  | Point 2 | White | 165.4574 | 66.5449 | 162.3053 | 64.1808
|  |  | Point 2 | Black | 0.2405 | 0.31667 | 0.27756 | 0.99788
|  |  | Point 3 | White | 172.4038 | 74.8922 | 158.9489 | 60.2406
|  |  | Point 3 | Black | 0.25807 | 0.35199 | 0.278 | 0.92871
|  |  | Point 4 | White | 170.9445 | 72.5281 | 159.7953 | 58.577
|  |  | Point 4 | Black | 0.24324 | 0.3444 | 0.26825 | 0.89018

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.073696 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12451 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16324 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21321 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.2733 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.32981 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.42495 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.54783 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.70018 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.88989 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1456 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4579 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8635 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3769 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 3.0208 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.8555 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.9383 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.2138 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.9124 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.5206 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.1182 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.5447 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 19.5315 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 23.6993 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.9007 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 37.8547 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 51.5139 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 65.2315 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 83.9399 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 105.4795 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 135.629 | PASS
31 | 148.75 | 175 | 201.25 | 172.3162 | PASS