## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S22475234
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 171.0203 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.17572 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 638.744 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1753 | 0.4963 | 0.03 | 0.015394 | PASS
Radius Red | 0.4497 | 0.5284 | 0.03 | 0.0013601 | PASS
Radius Green | 0.1252 | 0.5607 | 0.03 | 0.013264 | PASS
Radius Blue | 0.1198 | 0.303 | 0.03 | 0.020254 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.074063 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 58.8689 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 60.0119 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1753 | 0.4963 | 0.015394 | PASS | -0.0017 | 0.0153 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4497 | 0.5284 | 0.0013601 | PASS | -0.0013 | 0.0004 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1252 | 0.5607 | 0.013264 | PASS | 0.0132 | -0.0013 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1198 | 0.303 | 0.020254 | PASS | -0.0032 | 0.02 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 221.6995
Point 1 | 193.7098
Point 2 | 185.7711
Point 3 | 188.0477
Point 4 | 191.7543

## Curves
CURVES
Step | Luminance | Status
1 | 0.10072 | 0
2 | 0.045005 | 0
3 | 0.073053 | 0
4 | 0.073053 | 0
5 | 0.10291 | 0
6 | 0.14065 | 0
7 | 0.064619 | 0
8 | 0.10285 | 0
9 | 0.12209 | 0
10 | 0.12212 | 0
11 | 0.16035 | 0
12 | 0.16032 | 0
13 | 0.20965 | 0
14 | 0.20953 | 0
15 | 0.26825 | 0
16 | 0.26825 | 0
17 | 0.34703 | 0
18 | 0.34703 | 0
19 | 0.44597 | 0
20 | 0.44655 | 0
21 | 0.57497 | 0
22 | 0.58168 | 0
23 | 0.28077 | 0
24 | 0.43137 | 0
25 | 0.50668 | 0
26 | 0.54491 | 0
27 | 0.54491 | 0
28 | 0.69639 | 0
29 | 0.69522 | 0
30 | 0.88814 | 0
31 | 0.88814 | 0
32 | 1.1368 | 0
33 | 1.1397 | 0
34 | 1.4506 | 0
35 | 1.4497 | 0
36 | 1.86 | 0
37 | 1.8548 | 0
38 | 2.3766 | 0
39 | 2.3583 | 0
40 | 2.9887 | 0
41 | 3.0004 | 0
42 | 3.838 | 0
43 | 3.8147 | 0
44 | 4.9121 | 0
45 | 4.8741 | 0
46 | 6.2867 | 0
47 | 6.2546 | 0
48 | 8.0146 | 0
49 | 7.8803 | 0
50 | 10.2035 | 0
51 | 9.8708 | 0
52 | 12.5472 | 0
53 | 12.7632 | 0
54 | 16.2597 | 0
55 | 16.4699 | 0
56 | 21.0696 | 0
57 | 21.7672 | 0
58 | 10.8632 | 0
59 | 16.68 | 0
60 | 19.6016 | 0
61 | 19.3651 | 0
62 | 24.1284 | 0
63 | 27.6307 | 0
64 | 14.205 | 0
65 | 21.3615 | 0
66 | 24.6304 | 0
67 | 24.4552 | 0
68 | 32.6012 | 0
69 | 32.6012 | 0
70 | 41.3862 | 0
71 | 40.9485 | 0
72 | 50.8426 | 0
73 | 50.8134 | 0
74 | 64.4143 | 0
75 | 64.4435 | 0
76 | 82.0136 | 0
77 | 81.8093 | 0
78 | 103.1154 | 0
79 | 104.312 | 0
80 | 132.5644 | 0
81 | 132.4477 | 0
82 | 193.9725 | 0
83 | 95.2934 | 0
84 | 144.5892 | 0
85 | 169.0473 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 193.1553
150 | 200 | 171.0203 | PASS | Point 1 | 168.9014
|  |  |  | Point 2 | 161.9551
|  |  |  | Point 3 | 163.7062
Uniformity Test |  |  |  | Point 4 | 167.3837
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.17572 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 192.951 | White
100 | 638.744 | PASS | 0.30208 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.083648
0.05 | 0.1 | 0.074063 | PASS | Point 1 | 0.073199
|  |  |  | Point 2 | 0.070135
|  |  |  | Point 3 | 0.07104
|  |  |  | Point 4 | 0.072295

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 193.0677 | 75.0673 | 172.4621 | 64.0349
50 | 58.8689 | PASS | Point 1 | 169.0473 | 68.909 | 166.4206 | 66.1655
|  |  | Point 2 | 161.9551 | 67.1286 | 161.9842 | 63.7722
|  |  | Point 3 | 163.6479 | 75.7678 | 152.6446 | 59.0732
|  |  | Point 4 | 167.4713 | 73.8415 | 156.2053 | 58.8689

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 192.6007 | 74.9797 | 172.0827 | 63.9765
20 | 60.0119 | PASS | Point 0 | Black | 0.30325 | 0.37913 | 0.31288 | 1.0609
|  |  | Point 1 | White | 168.522 | 68.7922 | 166.012 | 66.0779
|  |  | Point 1 | Black | 0.2691 | 0.34294 | 0.30004 | 1.0784
|  |  | Point 2 | White | 161.6048 | 67.0702 | 161.5465 | 63.7138
|  |  | Point 2 | Black | 0.25121 | 0.33097 | 0.28944 | 1.0358
|  |  | Point 3 | White | 163.2684 | 75.651 | 152.3528 | 58.9564
|  |  | Point 3 | Black | 0.26087 | 0.38176 | 0.28054 | 0.98241
|  |  | Point 4 | White | 167.1794 | 73.7247 | 156.001 | 58.7813
|  |  | Point 4 | Black | 0.26253 | 0.37329 | 0.2848 | 0.97366

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.073141 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12191 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.15985 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20889 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26732 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34557 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.44422 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.54141 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.6923 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.88347 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1336 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4418 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.849 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3428 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 3.0033 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.8322 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8537 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1933 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.8395 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 10.081 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.4801 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.9153 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 19.3535 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 24.2159 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.2878 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 39.6643 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 50.2297 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 64.0349 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 81.2256 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 103.6115 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 131.6888 | PASS
31 | 148.75 | 175 | 201.25 | 168.3469 | PASS