## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23065548
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 178.6497 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.097726 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 715.4489 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1738 | 0.4949 | 0.03 | 0.014264 | PASS
Radius Red | 0.4509 | 0.5282 | 0.03 | 0.00022361 | PASS
Radius Green | 0.1252 | 0.5611 | 0.03 | 0.013231 | PASS
Radius Blue | 0.1206 | 0.299 | 0.03 | 0.016179 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.078389 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 57.9641 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 59.7176 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1738 | 0.4949 | 0.014264 | PASS | -0.0032 | 0.0139 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4509 | 0.5282 | 0.00022361 | PASS | -0.0001 | 0.0002 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1252 | 0.5611 | 0.013231 | PASS | 0.0132 | -0.0009 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1206 | 0.299 | 0.016179 | PASS | -0.0024 | 0.016 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 217.1464
Point 1 | 209.0035
Point 2 | 200.9772
Point 3 | 198.0878
Point 4 | 196.9787

## Curves
CURVES
Step | Luminance | Status
1 | 0.10557 | 0
2 | 0.047953 | 0
3 | 0.076906 | 0
4 | 0.076848 | 0
5 | 0.10793 | 0
6 | 0.14681 | 0
7 | 0.068442 | 0
8 | 0.10802 | 0
9 | 0.12757 | 0
10 | 0.12757 | 0
11 | 0.16715 | 0
12 | 0.16718 | 0
13 | 0.21817 | 0
14 | 0.23653 | 0
15 | 0.11237 | 0
16 | 0.17459 | 0
17 | 0.2055 | 0
18 | 0.20559 | 0
19 | 0.26291 | 0
20 | 0.26317 | 0
21 | 0.34031 | 0
22 | 0.34031 | 0
23 | 0.4375 | 0
24 | 0.43721 | 0
25 | 0.5633 | 0
26 | 0.56388 | 0
27 | 0.72207 | 0
28 | 0.72032 | 0
29 | 0.91791 | 0
30 | 0.91879 | 0
31 | 1.1806 | 0
32 | 1.2673 | 0
33 | 0.62079 | 0
34 | 0.94097 | 0
35 | 1.1044 | 0
36 | 1.1009 | 0
37 | 1.4071 | 0
38 | 1.4021 | 0
39 | 1.7912 | 0
40 | 1.7926 | 0
41 | 2.2809 | 0
42 | 2.2765 | 0
43 | 2.8991 | 0
44 | 2.9073 | 0
45 | 3.7271 | 0
46 | 3.7037 | 0
47 | 4.7107 | 0
48 | 4.7603 | 0
49 | 6.0153 | 0
50 | 6.0883 | 0
51 | 7.711 | 0
52 | 7.6118 | 0
53 | 9.722 | 0
54 | 9.5585 | 0
55 | 12.2583 | 0
56 | 12.41 | 0
57 | 15.46 | 0
58 | 15.9299 | 0
59 | 19.657 | 0
60 | 20.1386 | 0
61 | 24.1809 | 0
62 | 28.6756 | 0
63 | 14.1671 | 0
64 | 22.8879 | 0
65 | 26.2298 | 0
66 | 26.1918 | 0
67 | 30.3246 | 0
68 | 36.5413 | 0
69 | 17.4826 | 0
70 | 26.8339 | 0
71 | 31.6672 | 0
72 | 31.4045 | 0
73 | 36.1911 | 0
74 | 44.6259 | 0
75 | 22.4297 | 0
76 | 32.3677 | 0
77 | 36.4538 | 0
78 | 42.8747 | 0
79 | 40.8901 | 0
80 | 53.2651 | 0
81 | 53.3526 | 0
82 | 66.9243 | 0
83 | 67.1286 | 0
84 | 86.0997 | 0
85 | 86.3624 | 0
86 | 110.1201 | 0
87 | 109.1861 | 0
88 | 138.7811 | 0
89 | 137.9347 | 0
90 | 201.5026 | 0
91 | 99.5546 | 0
92 | 150.3973 | 0
93 | 175.731 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 189.828
150 | 200 | 178.6497 | PASS | Point 1 | 182.5898
|  |  |  | Point 2 | 175.5851
|  |  |  | Point 3 | 173.1042
Uniformity Test |  |  |  | Point 4 | 172.1411
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.097726 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 189.7697 | White
100 | 715.4489 | PASS | 0.26525 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.083385
0.05 | 0.1 | 0.078389 | PASS | Point 1 | 0.080146
|  |  |  | Point 2 | 0.077023
|  |  |  | Point 3 | 0.076176
|  |  |  | Point 4 | 0.075213

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 190.0615 | 73.0826 | 170.2148 | 63.0425
50 | 57.9641 | PASS | Point 1 | 182.7066 | 71.5941 | 173.4545 | 65.9028
|  |  | Point 2 | 175.6726 | 70.4267 | 167.0335 | 62.9841
|  |  | Point 3 | 173.221 | 74.5128 | 159.0072 | 59.044
|  |  | Point 4 | 172.1703 | 71.8276 | 159.6493 | 57.9641

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 189.5945 | 73.0242 | 170.0105 | 63.0425
20 | 59.7176 | PASS | Point 0 | Black | 0.26583 | 0.36804 | 0.28947 | 1.0116
|  |  | Point 1 | White | 182.4439 | 71.5649 | 173.3377 | 65.9028
|  |  | Point 1 | Black | 0.26031 | 0.3587 | 0.29537 | 1.0627
|  |  | Point 2 | White | 175.3516 | 70.3391 | 166.8292 | 62.955
|  |  | Point 2 | Black | 0.28217 | 0.35841 | 0.305 | 1.0542
|  |  | Point 3 | White | 172.8708 | 74.4252 | 158.9197 | 59.0148
|  |  | Point 3 | Black | 0.24271 | 0.37329 | 0.27065 | 0.92696
|  |  | Point 4 | White | 171.9368 | 71.7984 | 159.4742 | 57.9349
|  |  | Point 4 | Black | 0.25915 | 0.36395 | 0.2808 | 0.9217

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.077081 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12757 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16712 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20536 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26247 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33944 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.43634 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.56213 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.71886 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.91587 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1018 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4039 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.7906 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.2716 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9122 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7271 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.7224 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.0678 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.6527 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.5702 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.6231 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.8394 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 19.3009 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 25.9233 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 28.4917 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 39.1681 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 53.2651 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 66.9827 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 86.2165 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 110.0617 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 139.3648 | PASS
31 | 148.75 | 175 | 201.25 | 175.6143 | PASS