## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23065555
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 177.383 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.12547 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 667.0747 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1747 | 0.4894 | 0.03 | 0.0087092 | PASS
Radius Red | 0.4516 | 0.5278 | 0.03 | 0.00063246 | PASS
Radius Green | 0.1227 | 0.5618 | 0.03 | 0.010702 | PASS
Radius Blue | 0.1236 | 0.2913 | 0.03 | 0.0083217 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.072376 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 57.6722 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 60.3662 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1747 | 0.4894 | 0.0087092 | PASS | -0.0023 | 0.0084 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4516 | 0.5278 | 0.00063246 | PASS | 0.0006 | -0.0002 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1227 | 0.5618 | 0.010702 | PASS | 0.0107 | -0.0002 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1236 | 0.2913 | 0.0083217 | PASS | 0.0006 | 0.0083 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 203.2829
Point 1 | 187.5223
Point 2 | 179.2334
Point 3 | 189.5362
Point 4 | 184.4869

## Curves
CURVES
Step | Luminance | Status
1 | 0.09973 | 0
2 | 0.046027 | 0
3 | 0.07282 | 0
4 | 0.072762 | 0
5 | 0.10209 | 0
6 | 0.13852 | 0
7 | 0.065407 | 0
8 | 0.10209 | 0
9 | 0.12057 | 0
10 | 0.12965 | 0
11 | 0.12965 | 0
12 | 0.1696 | 0
13 | 0.16966 | 0
14 | 0.22132 | 0
15 | 0.22313 | 0
16 | 0.10726 | 0
17 | 0.16525 | 0
18 | 0.19441 | 0
19 | 0.20871 | 0
20 | 0.20871 | 0
21 | 0.26665 | 0
22 | 0.26673 | 0
23 | 0.34469 | 0
24 | 0.34527 | 0
25 | 0.44334 | 0
26 | 0.44334 | 0
27 | 0.57059 | 0
28 | 0.57059 | 0
29 | 0.2815 | 0
30 | 0.42554 | 0
31 | 0.49821 | 0
32 | 0.53557 | 0
33 | 0.53469 | 0
34 | 0.68413 | 0
35 | 0.68354 | 0
36 | 0.87034 | 0
37 | 0.87296 | 0
38 | 1.117 | 0
39 | 1.1181 | 0
40 | 1.4225 | 0
41 | 1.424 | 0
42 | 1.816 | 0
43 | 1.8177 | 0
44 | 2.3229 | 0
45 | 2.3288 | 0
46 | 2.9391 | 0
47 | 2.9391 | 0
48 | 3.7417 | 0
49 | 3.765 | 0
50 | 4.8187 | 0
51 | 4.7807 | 0
52 | 6.1116 | 0
53 | 6.0474 | 0
54 | 7.7723 | 0
55 | 7.854 | 0
56 | 9.7015 | 0
57 | 9.8562 | 0
58 | 12.8566 | 0
59 | 12.8624 | 0
60 | 15.7227 | 0
61 | 15.6702 | 0
62 | 20.6639 | 0
63 | 21.3557 | 0
64 | 26.3611 | 0
65 | 27.2951 | 0
66 | 31.4337 | 0
67 | 30.9083 | 0
68 | 41.2987 | 0
69 | 41.9992 | 0
70 | 53.7029 | 0
71 | 53.2943 | 0
72 | 67.4205 | 0
73 | 68.1209 | 0
74 | 85.2825 | 0
75 | 87.0045 | 0
76 | 109.7991 | 0
77 | 108.7483 | 0
78 | 139.5399 | 0
79 | 139.4524 | 0
80 | 190.1783 | 0
81 | 95.2642 | 0
82 | 142.6045 | 0
83 | 166.2455 | 0
84 | 178.2994 | 0
85 | 103.9618 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 190.8495
150 | 200 | 177.383 | PASS | Point 1 | 176.0521
|  |  |  | Point 2 | 168.3177
|  |  |  | Point 3 | 178.241
Uniformity Test |  |  |  | Point 4 | 173.4545
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.12547 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 190.8787 | White
100 | 667.0747 | PASS | 0.28614 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.078015
0.05 | 0.1 | 0.072376 | PASS | Point 1 | 0.071886
|  |  |  | Point 2 | 0.068646
|  |  |  | Point 3 | 0.072937
|  |  |  | Point 4 | 0.070397

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 190.9371 | 73.8123 | 171.5282 | 63.2468
50 | 57.6722 | PASS | Point 1 | 176.0813 | 71.0104 | 169.4851 | 65.5817
|  |  | Point 2 | 168.3761 | 68.8506 | 163.0058 | 62.6339
|  |  | Point 3 | 178.2994 | 75.7386 | 163.7938 | 59.6861
|  |  | Point 4 | 173.4253 | 72.6448 | 161.5173 | 57.6722

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 190.7328 | 73.462 | 171.4406 | 63.1009
20 | 60.3662 | PASS | Point 0 | Black | 0.28661 | 0.36629 | 0.30004 | 1.0408
|  |  | Point 1 | White | 176.2272 | 70.7769 | 169.456 | 65.465
|  |  | Point 1 | Black | 0.26037 | 0.34644 | 0.29741 | 1.0673
|  |  | Point 2 | White | 168.4344 | 68.6463 | 163.0641 | 62.5463
|  |  | Point 2 | Black | 0.26151 | 0.33915 | 0.29274 | 1.0361
|  |  | Point 3 | White | 178.1827 | 75.3008 | 163.6479 | 59.4526
|  |  | Point 3 | Black | 0.26682 | 0.37067 | 0.28302 | 0.96665
|  |  | Point 4 | White | 173.5129 | 72.2654 | 161.4297 | 57.4096
|  |  | Point 4 | Black | 0.25734 | 0.36074 | 0.28022 | 0.94272

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.073141 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12994 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16978 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20871 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26673 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.3444 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.44276 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.53411 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.68413 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.86946 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1158 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4217 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8186 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3101 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9332 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7563 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8041 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1554 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.8278 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.6957 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.2875 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.9066 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 21.306 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.5625 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 30.4997 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 43.7503 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 53.3235 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 68.0626 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 85.195 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 108.6608 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 140.0069 | PASS
31 | 148.75 | 175 | 201.25 | 178.1827 | PASS