## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23065560
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 179.4552 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.16167 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 575.1712 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1741 | 0.4933 | 0.03 | 0.012637 | PASS
Radius Red | 0.4497 | 0.5282 | 0.03 | 0.0013153 | PASS
Radius Green | 0.1237 | 0.5623 | 0.03 | 0.011704 | PASS
Radius Blue | 0.1213 | 0.299 | 0.03 | 0.01609 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.075085 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 56.9426 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 65.56 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1741 | 0.4933 | 0.012637 | PASS | -0.0029 | 0.0123 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4497 | 0.5282 | 0.0013153 | PASS | -0.0013 | 0.0002 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1237 | 0.5623 | 0.011704 | PASS | 0.0117 | 0.0003 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1213 | 0.299 | 0.01609 | PASS | -0.0017 | 0.016 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 196.2782
Point 1 | 181.364
Point 2 | 173.5712
Point 3 | 179.7004
Point 4 | 166.8584

## Curves
CURVES
Step | Luminance | Status
1 | 0.094418 | 0
2 | 0.043429 | 0
3 | 0.069055 | 0
4 | 0.08178 | 0
5 | 0.075447 | 0
6 | 0.075418 | 0
7 | 0.10554 | 0
8 | 0.13131 | 0
9 | 0.13134 | 0
10 | 0.17156 | 0
11 | 0.16198 | 0
12 | 0.16207 | 0
13 | 0.21125 | 0
14 | 0.21128 | 0
15 | 0.26997 | 0
16 | 0.26994 | 0
17 | 0.34878 | 0
18 | 0.34878 | 0
19 | 0.44772 | 0
20 | 0.41912 | 0
21 | 0.4197 | 0
22 | 0.54024 | 0
23 | 0.54082 | 0
24 | 0.68967 | 0
25 | 0.69113 | 0
26 | 0.8788 | 0
27 | 0.87793 | 0
28 | 1.1272 | 0
29 | 1.1281 | 0
30 | 1.4333 | 0
31 | 1.4339 | 0
32 | 1.837 | 0
33 | 1.8399 | 0
34 | 2.3515 | 0
35 | 2.3396 | 0
36 | 2.9712 | 0
37 | 2.9624 | 0
38 | 3.8117 | 0
39 | 3.7826 | 0
40 | 4.8537 | 0
41 | 4.8391 | 0
42 | 6.1525 | 0
43 | 6.1642 | 0
44 | 7.9504 | 0
45 | 7.8745 | 0
46 | 9.9992 | 0
47 | 9.9263 | 0
48 | 12.6931 | 0
49 | 12.9179 | 0
50 | 16.0554 | 0
51 | 16.2568 | 0
52 | 20.8332 | 0
53 | 20.2174 | 0
54 | 26.758 | 0
55 | 25.5877 | 0
56 | 30.6457 | 0
57 | 30.383 | 0
58 | 42.9623 | 0
59 | 42.1451 | 0
60 | 38.6719 | 0
61 | 35.6657 | 0
62 | 30.2371 | 0
63 | 32.7471 | 0
64 | 31.3753 | 0
65 | 42.6121 | 0
66 | 42.5829 | 0
67 | 54.0239 | 0
68 | 54.4617 | 0
69 | 68.7339 | 0
70 | 69.5803 | 0
71 | 87.0921 | 0
72 | 86.9753 | 0
73 | 111.3751 | 0
74 | 111.4627 | 0
75 | 142.5461 | 0
76 | 141.4954 | 0
77 | 179.8171 | 0
78 | 50.8426 | 0
79 | 50.4632 | 0
80 | 63.8597 | 0
81 | 64.1224 | 0
82 | 81.4299 | 0
83 | 82.0428 | 0
84 | 103.845 | 0
85 | 103.9618 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 196.1031
150 | 200 | 179.4552 | PASS | Point 1 | 181.2472
|  |  |  | Point 2 | 173.5129
|  |  |  | Point 3 | 179.642
Uniformity Test |  |  |  | Point 4 | 166.7708
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.16167 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 196.0739 | White
100 | 575.1712 | PASS | 0.3409 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.082189
0.05 | 0.1 | 0.075085 | PASS | Point 1 | 0.075826
|  |  |  | Point 2 | 0.072528
|  |  |  | Point 3 | 0.075388
|  |  |  | Point 4 | 0.069493

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 196.4825 | 74.5711 | 175.6143 | 63.3928
50 | 56.9426 | PASS | Point 1 | 181.3932 | 71.7109 | 175.0014 | 66.253
|  |  | Point 2 | 173.542 | 70.8061 | 169.0765 | 63.8305
|  |  | Point 3 | 179.7295 | 76.8185 | 164.4651 | 59.8612
|  |  | Point 4 | 166.8584 | 73.3161 | 158.5403 | 56.9426

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 196.0739 | 74.5711 | 175.5559 | 63.4219
20 | 65.56 | PASS | Point 0 | Black | 0.34119 | 0.39226 | 0.3409 | 0.95556
|  |  | Point 1 | White | 181.0721 | 71.7401 | 174.8554 | 66.3406
|  |  | Point 1 | Black | 0.35987 | 0.39139 | 0.36687 | 0.98037
|  |  | Point 2 | White | 173.2794 | 70.8645 | 168.9306 | 63.8889
|  |  | Point 2 | Black | 0.3266 | 0.37359 | 0.35403 | 0.93221
|  |  | Point 3 | White | 179.5252 | 76.8768 | 164.4067 | 59.8904
|  |  | Point 3 | Black | 0.31229 | 0.40598 | 0.32368 | 0.89223
|  |  | Point 4 | White | 166.7416 | 73.3745 | 158.4819 | 57.0593
|  |  | Point 4 | Black | 0.27091 | 0.38351 | 0.29683 | 0.87034

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.075651 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.13146 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.1621 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21128 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26997 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34878 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.41941 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.53995 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.68909 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.87997 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1281 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4333 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8399 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.339 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9595 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7884 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8128 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1612 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.8453 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 10.0197 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.6786 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.8336 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.8303 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.796 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 34.8777 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 44.3924 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 53.7321 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 68.9382 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 86.7418 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 111.3751 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 142.3418 | PASS
31 | 148.75 | 175 | 201.25 | 179.7587 | PASS