## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23166791
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 173.5771 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.16758 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 688.1286 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1746 | 0.4925 | 0.03 | 0.011748 | PASS
Radius Red | 0.4512 | 0.528 | 0.03 | 0.0002 | PASS
Radius Green | 0.1232 | 0.5623 | 0.03 | 0.011204 | PASS
Radius Blue | 0.122 | 0.2958 | 0.03 | 0.012839 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.078943 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 59.044 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 67.1911 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1746 | 0.4925 | 0.011748 | PASS | -0.0024 | 0.0115 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4512 | 0.528 | 0.0002 | PASS | 0.0002 | 0 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1232 | 0.5623 | 0.011204 | PASS | 0.0112 | 0.0003 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.122 | 0.2958 | 0.012839 | PASS | -0.001 | 0.0128 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 193.6806
Point 1 | 163.6479
Point 2 | 165.8077
Point 3 | 171.2363
Point 4 | 173.0751

## Curves
CURVES
Step | Luminance | Status
1 | 0.090915 | 0
2 | 0.04127 | 0
3 | 0.066078 | 0
4 | 0.07857 | 0
5 | 0.07854 | 0
6 | 0.10983 | 0
7 | 0.12638 | 0
8 | 0.12641 | 0
9 | 0.16537 | 0
10 | 0.16537 | 0
11 | 0.21566 | 0
12 | 0.21563 | 0
13 | 0.27526 | 0
14 | 0.27529 | 0
15 | 0.35578 | 0
16 | 0.33389 | 0
17 | 0.33418 | 0
18 | 0.42875 | 0
19 | 0.42904 | 0
20 | 0.55162 | 0
21 | 0.55221 | 0
22 | 0.70485 | 0
23 | 0.70485 | 0
24 | 0.89952 | 0
25 | 0.90069 | 0
26 | 1.1534 | 0
27 | 1.1511 | 0
28 | 1.4669 | 0
29 | 1.4692 | 0
30 | 1.8767 | 0
31 | 1.8778 | 0
32 | 2.3866 | 0
33 | 2.3895 | 0
34 | 3.0266 | 0
35 | 3.0412 | 0
36 | 3.8818 | 0
37 | 3.8614 | 0
38 | 4.9558 | 0
39 | 4.9617 | 0
40 | 6.2371 | 0
41 | 6.3334 | 0
42 | 8.0467 | 0
43 | 7.4688 | 0
44 | 7.5272 | 0
45 | 9.5819 | 0
46 | 9.6315 | 0
47 | 12.2378 | 0
48 | 12.1736 | 0
49 | 15.498 | 0
50 | 15.6556 | 0
51 | 19.4002 | 0
52 | 19.4031 | 0
53 | 25.862 | 0
54 | 25.252 | 0
55 | 33.4184 | 0
56 | 30.2662 | 0
57 | 41.6489 | 0
58 | 42.6413 | 0
59 | 51.9225 | 0
60 | 51.9809 | 0
61 | 65.465 | 0
62 | 66.1071 | 0
63 | 84.4653 | 0
64 | 83.473 | 0
65 | 106.1508 | 0
66 | 106.0048 | 0
67 | 136.4462 | 0
68 | 134.7826 | 0
69 | 173.2794 | 0
70 | 33.2724 | 0
71 | 36.6289 | 0
72 | 45.7058 | 0
73 | 45.7058 | 0
74 | 23.5651 | 0
75 | 35.3739 | 0
76 | 40.248 | 0
77 | 36.9499 | 0
78 | 50.8426 | 0
79 | 50.4632 | 0
80 | 63.8597 | 0
81 | 64.1224 | 0
82 | 81.4299 | 0
83 | 82.0428 | 0
84 | 103.845 | 0
85 | 103.9618 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 193.8557
150 | 200 | 173.5771 | PASS | Point 1 | 163.8814
|  |  |  | Point 2 | 165.9244
|  |  |  | Point 3 | 171.1196
Uniformity Test |  |  |  | Point 4 | 173.1042
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.16758 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 193.7098 | White
100 | 688.1286 | PASS | 0.2815 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.08826
0.05 | 0.1 | 0.078943 | PASS | Point 1 | 0.074542
|  |  |  | Point 2 | 0.075534
|  |  |  | Point 3 | 0.077928
|  |  |  | Point 4 | 0.078453

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 194.1184 | 76.906 | 176.2564 | 65.9612
50 | 59.044 | PASS | Point 1 | 163.9689 | 69.6094 | 161.9259 | 66.1655
|  |  | Point 2 | 166.012 | 69.9597 | 162.5096 | 66.0779
|  |  | Point 3 | 171.2655 | 75.797 | 158.5694 | 59.044
|  |  | Point 4 | 173.1626 | 76.0013 | 161.7799 | 59.832

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 193.6514 | 76.8477 | 176.198 | 65.9612
20 | 67.1911 | PASS | Point 0 | Black | 0.28194 | 0.37709 | 0.3015 | 0.97336
|  |  | Point 1 | White | 163.7062 | 69.6386 | 161.7216 | 66.1946
|  |  | Point 1 | Black | 0.23854 | 0.33243 | 0.27243 | 0.95819
|  |  | Point 2 | White | 165.7785 | 69.9597 | 162.3637 | 66.1071
|  |  | Point 2 | Black | 0.23819 | 0.33039 | 0.27029 | 0.96607
|  |  | Point 3 | White | 171.0904 | 75.7678 | 158.5403 | 59.1024
|  |  | Point 3 | Black | 0.248 | 0.37067 | 0.27 | 0.868
|  |  | Point 4 | White | 173.0167 | 76.0013 | 161.8091 | 59.832
|  |  | Point 4 | Black | 0.25871 | 0.37213 | 0.28001 | 0.89048

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.078774 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12667 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16552 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21572 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.27529 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33389 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.42904 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.55191 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.70631 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.89952 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1514 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4649 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8708 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3854 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 3.05 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.8789 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8916 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.2809 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.5213 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.5673 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.2291 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.6439 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.0744 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 25.7599 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.4045 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 40.3355 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 51.8058 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 66.0779 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 83.3854 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 107.4641 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 135.8041 | PASS
31 | 148.75 | 175 | 201.25 | 173.3086 | PASS