## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23166792
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 171.2538 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.14799 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 682.7403 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1728 | 0.4948 | 0.03 | 0.014425 | PASS
Radius Red | 0.4502 | 0.5281 | 0.03 | 0.00080623 | PASS
Radius Green | 0.1238 | 0.5614 | 0.03 | 0.011815 | PASS
Radius Blue | 0.1215 | 0.2978 | 0.03 | 0.014876 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.075295 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 59.2775 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 65.0599 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1728 | 0.4948 | 0.014425 | PASS | -0.0042 | 0.0138 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4502 | 0.5281 | 0.00080623 | PASS | -0.0008 | 0.0001 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1238 | 0.5614 | 0.011815 | PASS | 0.0118 | -0.0006 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1215 | 0.2978 | 0.014876 | PASS | -0.0015 | 0.0148 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 214.5489
Point 1 | 198.7882
Point 2 | 184.6037
Point 3 | 188.7773
Point 4 | 191.6668

## Curves
CURVES
Step | Luminance | Status
1 | 0.10393 | 0
2 | 0.04807 | 0
3 | 0.076089 | 0
4 | 0.07606 | 0
5 | 0.10685 | 0
6 | 0.14471 | 0
7 | 0.068442 | 0
8 | 0.10688 | 0
9 | 0.12614 | 0
10 | 0.12611 | 0
11 | 0.1654 | 0
12 | 0.16537 | 0
13 | 0.21583 | 0
14 | 0.21595 | 0
15 | 0.27578 | 0
16 | 0.27581 | 0
17 | 0.35695 | 0
18 | 0.38263 | 0
19 | 0.18776 | 0
20 | 0.28515 | 0
21 | 0.33448 | 0
22 | 0.33448 | 0
23 | 0.42991 | 0
24 | 0.42991 | 0
25 | 0.55367 | 0
26 | 0.55367 | 0
27 | 0.7066 | 0
28 | 0.70806 | 0
29 | 0.90332 | 0
30 | 0.90069 | 0
31 | 1.1593 | 0
32 | 1.1596 | 0
33 | 1.4689 | 0
34 | 1.4695 | 0
35 | 1.8741 | 0
36 | 1.8881 | 0
37 | 2.4055 | 0
38 | 2.3889 | 0
39 | 3.0412 | 0
40 | 3.0325 | 0
41 | 3.9168 | 0
42 | 4.1999 | 0
43 | 2.1005 | 0
44 | 3.1229 | 0
45 | 3.6687 | 0
46 | 3.6921 | 0
47 | 4.629 | 0
48 | 4.6727 | 0
49 | 5.9569 | 0
50 | 6.0066 | 0
51 | 7.6235 | 0
52 | 7.46 | 0
53 | 9.4564 | 0
54 | 9.7453 | 0
55 | 12.3225 | 0
56 | 12.0394 | 0
57 | 15.1827 | 0
58 | 15.5505 | 0
59 | 19.0616 | 0
60 | 19.9138 | 0
61 | 24.2743 | 0
62 | 25.9846 | 0
63 | 29.7993 | 0
64 | 36.1911 | 0
65 | 17.039 | 0
66 | 26.6355 | 0
67 | 28.8916 | 0
68 | 32.4552 | 0
69 | 32.2217 | 0
70 | 41.4738 | 0
71 | 43.0207 | 0
72 | 52.7397 | 0
73 | 52.6814 | 0
74 | 67.041 | 0
75 | 66.7784 | 0
76 | 84.5529 | 0
77 | 85.6619 | 0
78 | 108.4565 | 0
79 | 107.3766 | 0
80 | 137.5553 | 0
81 | 137.3802 | 0
82 | 199.3136 | 0
83 | 99.9633 | 0
84 | 149.58 | 0
85 | 174.2717 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 187.6682
150 | 200 | 171.2538 | PASS | Point 1 | 174.1841
|  |  |  | Point 2 | 161.8091
|  |  |  | Point 3 | 164.9613
Uniformity Test |  |  |  | Point 4 | 167.6464
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.14799 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 187.6099 | White
100 | 682.7403 | PASS | 0.27479 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.082627
0.05 | 0.1 | 0.075295 | PASS | Point 1 | 0.076585
|  |  |  | Point 2 | 0.071186
|  |  |  | Point 3 | 0.072674
|  |  |  | Point 4 | 0.073404

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 188.0185 | 73.608 | 169.4851 | 64.4435
50 | 59.2775 | PASS | Point 1 | 174.5636 | 69.6678 | 167.2962 | 66.4865
|  |  | Point 2 | 162.1302 | 66.4573 | 159.0948 | 63.6262
|  |  | Point 3 | 165.1364 | 74.8338 | 154.7752 | 59.511
|  |  | Point 4 | 167.8799 | 72.7616 | 157.2561 | 59.2775

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 187.6682 | 73.5788 | 169.3976 | 64.4726
20 | 65.0599 | PASS | Point 0 | Black | 0.27631 | 0.35461 | 0.28997 | 0.98854
|  |  | Point 1 | White | 174.3301 | 69.697 | 167.267 | 66.5741
|  |  | Point 1 | Black | 0.26752 | 0.33623 | 0.28851 | 1.0233
|  |  | Point 2 | White | 161.9551 | 66.4865 | 159.0656 | 63.6846
|  |  | Point 2 | Black | 0.23816 | 0.31521 | 0.27076 | 0.97366
|  |  | Point 3 | White | 164.9321 | 74.8338 | 154.6585 | 59.511
|  |  | Point 3 | Black | 0.24178 | 0.3587 | 0.26376 | 0.90098
|  |  | Point 4 | White | 167.7048 | 72.7908 | 157.2561 | 59.3067
|  |  | Point 4 | Black | 0.25465 | 0.3514 | 0.27461 | 0.90478

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.076906 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12705 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16622 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21685 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.27677 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33564 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.43108 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.55512 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.70806 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.90244 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1619 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.476 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8945 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3927 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 3.0471 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.6775 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.7078 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 5.9394 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.4659 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.6432 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.5472 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.2324 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 19.8292 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 24.966 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 33.1849 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 38.8178 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 52.7105 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 67.1286 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 85.8371 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 107.3474 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 137.0007 | PASS
31 | 148.75 | 175 | 201.25 | 174.6511 | PASS