## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S23349996
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 178.1885 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.14402 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 692.9126 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1727 | 0.4957 | 0.03 | 0.015316 | PASS
Radius Red | 0.4515 | 0.5283 | 0.03 | 0.0005831 | PASS
Radius Green | 0.1222 | 0.5629 | 0.03 | 0.01024 | PASS
Radius Blue | 0.1193 | 0.3049 | 0.03 | 0.02221 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.0747 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 61.5248 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 67.1338 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1727 | 0.4957 | 0.015316 | PASS | -0.0043 | 0.0147 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4515 | 0.5283 | 0.0005831 | PASS | 0.0005 | 0.0003 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1222 | 0.5629 | 0.01024 | PASS | 0.0102 | 0.0009 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1193 | 0.3049 | 0.02221 | PASS | -0.0037 | 0.0219 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 195.3442
Point 1 | 178.1243
Point 2 | 172.0244
Point 3 | 175.8186
Point 4 | 169.0473

## Curves
CURVES
Step | Luminance | Status
1 | 0.093426 | 0
2 | 0.042525 | 0
3 | 0.067975 | 0
4 | 0.080729 | 0
5 | 0.074396 | 0
6 | 0.074425 | 0
7 | 0.10425 | 0
8 | 0.13008 | 0
9 | 0.13011 | 0
10 | 0.17024 | 0
11 | 0.17024 | 0
12 | 0.22193 | 0
13 | 0.20944 | 0
14 | 0.20965 | 0
15 | 0.26805 | 0
16 | 0.26805 | 0
17 | 0.34615 | 0
18 | 0.34615 | 0
19 | 0.4448 | 0
20 | 0.44422 | 0
21 | 0.57234 | 0
22 | 0.53615 | 0
23 | 0.53674 | 0
24 | 0.68384 | 0
25 | 0.68559 | 0
26 | 0.87238 | 0
27 | 0.87296 | 0
28 | 1.1222 | 0
29 | 1.1216 | 0
30 | 1.4275 | 0
31 | 1.4269 | 0
32 | 1.8206 | 0
33 | 1.8209 | 0
34 | 2.3276 | 0
35 | 2.3346 | 0
36 | 2.9595 | 0
37 | 2.9361 | 0
38 | 3.765 | 0
39 | 3.7446 | 0
40 | 4.7895 | 0
41 | 4.7953 | 0
42 | 6.062 | 0
43 | 6.1262 | 0
44 | 7.819 | 0
45 | 7.7315 | 0
46 | 9.7103 | 0
47 | 9.8679 | 0
48 | 12.3283 | 0
49 | 12.7252 | 0
50 | 16.0204 | 0
51 | 16.5049 | 0
52 | 19.9635 | 0
53 | 19.9576 | 0
54 | 26.6355 | 0
55 | 27.1842 | 0
56 | 32.3093 | 0
57 | 34.0021 | 0
58 | 41.2403 | 0
59 | 43.6628 | 0
60 | 53.3526 | 0
61 | 52.944 | 0
62 | 68.0334 | 0
63 | 68.5296 | 0
64 | 85.8954 | 0
65 | 86.8878 | 0
66 | 110.8206 | 0
67 | 110.5579 | 0
68 | 140.182 | 0
69 | 139.4232 | 0
70 | 178.3578 | 0
71 | 36.6289 | 0
72 | 45.7058 | 0
73 | 45.7058 | 0
74 | 23.5651 | 0
75 | 35.3739 | 0
76 | 40.248 | 0
77 | 36.9499 | 0
78 | 50.8426 | 0
79 | 50.4632 | 0
80 | 63.8597 | 0
81 | 64.1224 | 0
82 | 81.4299 | 0
83 | 82.0428 | 0
84 | 103.845 | 0
85 | 103.9618 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 195.5194
150 | 200 | 178.1885 | PASS | Point 1 | 178.0951
|  |  |  | Point 2 | 172.0827
|  |  |  | Point 3 | 175.9937
Uniformity Test |  |  |  | Point 4 | 169.2516
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.14402 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 195.461 | White
100 | 692.9126 | PASS | 0.28209 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.082043
0.05 | 0.1 | 0.0747 | PASS | Point 1 | 0.074717
|  |  |  | Point 2 | 0.072178
|  |  |  | Point 3 | 0.073871
|  |  |  | Point 4 | 0.070689

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 195.8112 | 78.3654 | 176.8401 | 67.7999
50 | 61.5248 | PASS | Point 1 | 178.4745 | 74.7462 | 172.6081 | 70.4559
|  |  | Point 2 | 172.3454 | 72.7908 | 168.4636 | 68.4128
|  |  | Point 3 | 176.198 | 80.1165 | 164.0857 | 63.9181
|  |  | Point 4 | 169.4851 | 77.2855 | 161.1962 | 61.5248

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 195.461 | 78.3945 | 176.6942 | 67.8583
20 | 67.1338 | PASS | Point 0 | Black | 0.28331 | 0.35345 | 0.29828 | 1.0063
|  |  | Point 1 | White | 178.1535 | 74.7171 | 172.5789 | 70.485
|  |  | Point 1 | Black | 0.26402 | 0.32689 | 0.289 | 0.9938
|  |  | Point 2 | White | 172.1119 | 72.7616 | 168.3761 | 68.4712
|  |  | Point 2 | Black | 0.25205 | 0.31871 | 0.28156 | 0.98504
|  |  | Point 3 | White | 175.9353 | 80.1165 | 163.9397 | 63.8597
|  |  | Point 3 | Black | 0.25804 | 0.35637 | 0.27926 | 0.9398
|  |  | Point 4 | White | 169.2225 | 77.2855 | 161.167 | 61.5248
|  |  | Point 4 | Black | 0.24902 | 0.34994 | 0.27496 | 0.91645

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.074951 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.13067 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.17068 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20994 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26843 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34644 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.44509 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.53645 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.6853 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.87267 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1202 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4249 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8183 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3262 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9478 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7475 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.807 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1175 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.7315 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 10.0226 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.6902 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.9941 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.696 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.6355 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 32.426 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 42.0867 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 53.2359 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 67.3621 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 86.8294 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 109.8574 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 139.4524 | PASS
31 | 148.75 | 175 | 201.25 | 178.4162 | PASS