## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24017774
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 173.8631 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.1545 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 796.6449 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1722 | 0.4942 | 0.03 | 0.014046 | PASS
Radius Red | 0.4506 | 0.5282 | 0.03 | 0.00044721 | PASS
Radius Green | 0.1232 | 0.5605 | 0.03 | 0.0113 | PASS
Radius Blue | 0.1197 | 0.3006 | 0.03 | 0.017907 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.071553 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 57.9349 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 80.4809 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1722 | 0.4942 | 0.014046 | PASS | -0.0048 | 0.0132 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4506 | 0.5282 | 0.00044721 | PASS | -0.0004 | 0.0002 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1232 | 0.5605 | 0.0113 | PASS | 0.0112 | -0.0015 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1197 | 0.3006 | 0.017907 | PASS | -0.0033 | 0.0176 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 204.6547
Point 1 | 185.4793
Point 2 | 176.8985
Point 3 | 181.3932
Point 4 | 175.3808

## Curves
CURVES
Step | Luminance | Status
1 | 0.097541 | 0
2 | 0.045589 | 0
3 | 0.071769 | 0
4 | 0.071682 | 0
5 | 0.10055 | 0
6 | 0.13566 | 0
7 | 0.064852 | 0
8 | 0.10061 | 0
9 | 0.11841 | 0
10 | 0.12716 | 0
11 | 0.12719 | 0
12 | 0.16636 | 0
13 | 0.16642 | 0
14 | 0.21706 | 0
15 | 0.21884 | 0
16 | 0.10636 | 0
17 | 0.16292 | 0
18 | 0.19091 | 0
19 | 0.20509 | 0
20 | 0.20492 | 0
21 | 0.26204 | 0
22 | 0.26206 | 0
23 | 0.33885 | 0
24 | 0.33856 | 0
25 | 0.43517 | 0
26 | 0.43488 | 0
27 | 0.55979 | 0
28 | 0.56009 | 0
29 | 0.71711 | 0
30 | 0.71623 | 0
31 | 0.91324 | 0
32 | 0.91237 | 0
33 | 1.1704 | 0
34 | 1.1695 | 0
35 | 1.1736 | 0
36 | 1.4932 | 0
37 | 1.4891 | 0
38 | 0.74863 | 0
39 | 1.1161 | 0
40 | 1.3032 | 0
41 | 1.3966 | 0
42 | 1.4004 | 0
43 | 1.7833 | 0
44 | 1.7824 | 0
45 | 2.2841 | 0
46 | 2.278 | 0
47 | 2.8827 | 0
48 | 2.8924 | 0
49 | 3.6775 | 0
50 | 3.7096 | 0
51 | 4.699 | 0
52 | 4.7282 | 0
53 | 6.0591 | 0
54 | 5.9453 | 0
55 | 7.5739 | 0
56 | 7.6614 | 0
57 | 9.5469 | 0
58 | 9.684 | 0
59 | 12.3458 | 0
60 | 12.5793 | 0
61 | 15.8832 | 0
62 | 15.8511 | 0
63 | 20.2816 | 0
64 | 19.5549 | 0
65 | 25.2433 | 0
66 | 26.0897 | 0
67 | 30.4706 | 0
68 | 30.9375 | 0
69 | 33.8854 | 0
70 | 42.4078 | 0
71 | 41.8824 | 0
72 | 52.8273 | 0
73 | 52.0393 | 0
74 | 66.5741 | 0
75 | 67.0119 | 0
76 | 84.2902 | 0
77 | 84.3486 | 0
78 | 108.2814 | 0
79 | 108.8359 | 0
80 | 136.4754 | 0
81 | 137.2342 | 0
82 | 186.384 | 0
83 | 94.2135 | 0
84 | 140.4155 | 0
85 | 163.0641 | 0
86 | 174.6803 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 192.7175
150 | 200 | 173.8631 | PASS | Point 1 | 174.5928
|  |  |  | Point 2 | 166.4498
|  |  |  | Point 3 | 170.4775
Uniformity Test |  |  |  | Point 4 | 165.078
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.1545 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 192.6591 | White
100 | 796.6449 | PASS | 0.24184 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.079387
0.05 | 0.1 | 0.071553 | PASS | Point 1 | 0.071886
|  |  |  | Point 2 | 0.0685
|  |  |  | Point 3 | 0.07031
|  |  |  | Point 4 | 0.067683

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 192.951 | 75.7678 | 174.5052 | 64.064
50 | 57.9349 | PASS | Point 1 | 174.7095 | 70.8645 | 168.055 | 66.4281
|  |  | Point 2 | 166.5373 | 69.3468 | 163.7938 | 64.8521
|  |  | Point 3 | 170.6818 | 76.2056 | 160.2331 | 59.6861
|  |  | Point 4 | 165.078 | 74.6295 | 157.3144 | 57.9349

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 192.5715 | 75.7094 | 174.3884 | 64.0057
20 | 80.4809 | PASS | Point 0 | Black | 0.24181 | 0.32251 | 0.25599 | 0.79445
|  |  | Point 1 | White | 174.5636 | 70.8645 | 167.8507 | 66.4281
|  |  | Point 1 | Black | 0.19961 | 0.29303 | 0.23031 | 0.82539
|  |  | Point 2 | White | 166.4498 | 69.3176 | 163.5895 | 64.8521
|  |  | Point 2 | Black | 0.20714 | 0.28206 | 0.23089 | 0.79445
|  |  | Point 3 | White | 170.5067 | 76.1472 | 160.1163 | 59.6277
|  |  | Point 3 | Black | 0.22479 | 0.32543 | 0.24347 | 0.73871
|  |  | Point 4 | White | 165.0196 | 74.6295 | 157.1393 | 57.8473
|  |  | Point 4 | Black | 0.24662 | 0.31871 | 0.25226 | 0.7101

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.07209 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.1274 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16645 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20483 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26186 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33856 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.43488 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.56067 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.71594 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.91207 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1707 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.3963 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.7807 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.2722 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.8994 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.6775 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.7224 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.0036 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.7344 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.6461 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.3312 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.5038 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.2641 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 25.897 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 30.9083 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 39.3432 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 52.0684 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 66.72 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 84.4069 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 108.9818 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 138.0223 | PASS
31 | 148.75 | 175 | 201.25 | 174.6219 | PASS