## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24090880
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 177.6048 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.12508 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 921.8359 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.171 | 0.4983 | 0.03 | 0.018311 | PASS
Radius Red | 0.4508 | 0.5285 | 0.03 | 0.00053852 | PASS
Radius Green | 0.1216 | 0.5615 | 0.03 | 0.009613 | PASS
Radius Blue | 0.1172 | 0.3093 | 0.03 | 0.026932 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.077752 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 62.0502 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 83.0933 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.171 | 0.4983 | 0.018311 | PASS | -0.006 | 0.0173 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4508 | 0.5285 | 0.00053852 | PASS | -0.0002 | 0.0005 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1216 | 0.5615 | 0.009613 | PASS | 0.0096 | -0.0005 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1172 | 0.3093 | 0.026932 | PASS | -0.0058 | 0.0263 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 218.6058
Point 1 | 202.3782
Point 2 | 193.5055
Point 3 | 196.9203
Point 4 | 199.0509

## Curves
CURVES
Step | Luminance | Status
1 | 0.10545 | 0
2 | 0.048508 | 0
3 | 0.07714 | 0
4 | 0.07714 | 0
5 | 0.10837 | 0
6 | 0.14684 | 0
7 | 0.069259 | 0
8 | 0.10846 | 0
9 | 0.12798 | 0
10 | 0.12801 | 0
11 | 0.16785 | 0
12 | 0.16785 | 0
13 | 0.2191 | 0
14 | 0.23699 | 0
15 | 0.114 | 0
16 | 0.17567 | 0
17 | 0.20667 | 0
18 | 0.20673 | 0
19 | 0.26463 | 0
20 | 0.2646 | 0
21 | 0.34206 | 0
22 | 0.34206 | 0
23 | 0.43955 | 0
24 | 0.44013 | 0
25 | 0.56709 | 0
26 | 0.56651 | 0
27 | 0.7247 | 0
28 | 0.77577 | 0
29 | 0.3838 | 0
30 | 0.58081 | 0
31 | 0.67712 | 0
32 | 0.67712 | 0
33 | 0.86275 | 0
34 | 0.8645 | 0
35 | 1.1088 | 0
36 | 1.1082 | 0
37 | 1.4103 | 0
38 | 1.4152 | 0
39 | 1.8107 | 0
40 | 1.8101 | 0
41 | 2.3011 | 0
42 | 2.3078 | 0
43 | 2.9274 | 0
44 | 2.9154 | 0
45 | 3.7417 | 0
46 | 3.7096 | 0
47 | 4.7428 | 0
48 | 4.7895 | 0
49 | 6.0649 | 0
50 | 6.0649 | 0
51 | 7.784 | 0
52 | 7.7752 | 0
53 | 9.7278 | 0
54 | 9.6782 | 0
55 | 12.413 | 0
56 | 12.7136 | 0
57 | 15.5826 | 0
58 | 15.6935 | 0
59 | 19.7796 | 0
60 | 20.4626 | 0
61 | 27.0266 | 0
62 | 29.1776 | 0
63 | 14.4414 | 0
64 | 22.4151 | 0
65 | 26.402 | 0
66 | 26.4049 | 0
67 | 32.8638 | 0
68 | 31.0543 | 0
69 | 43.2542 | 0
70 | 43.5168 | 0
71 | 53.09 | 0
72 | 52.8857 | 0
73 | 68.0626 | 0
74 | 67.5372 | 0
75 | 85.0198 | 0
76 | 85.2241 | 0
77 | 110.1785 | 0
78 | 108.1938 | 0
79 | 139.4232 | 0
80 | 138.46 | 0
81 | 202.6992 | 0
82 | 101.335 | 0
83 | 151.9441 | 0
84 | 177.0736 | 0
85 | 103.9618 | 0
86 | 132.506 | 0
87 | 133.2065 | 0
88 | 206.0556 | 0
89 | 102.269 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 193.3888
150 | 200 | 177.6048 | PASS | Point 1 | 176.9568
|  |  |  | Point 2 | 170.6234
|  |  |  | Point 3 | 172.6081
Uniformity Test |  |  |  | Point 4 | 174.4468
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.12508 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 193.4471 | White
100 | 921.8359 | PASS | 0.20985 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.084845
0.05 | 0.1 | 0.077752 | PASS | Point 1 | 0.077373
|  |  |  | Point 2 | 0.074659
|  |  |  | Point 3 | 0.075797
|  |  |  | Point 4 | 0.076089

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 193.5639 | 81.8677 | 175.8769 | 69.0841
50 | 62.0502 | PASS | Point 1 | 176.519 | 76.6434 | 172.1995 | 71.6817
|  |  | Point 2 | 170.7402 | 76.0304 | 169.1057 | 70.7185
|  |  | Point 3 | 172.9291 | 82.3639 | 161.1962 | 62.0502
|  |  | Point 4 | 174.7095 | 81.3424 | 163.9105 | 62.1086

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 193.1261 | 81.8969 | 175.8478 | 69.1133
20 | 83.0933 | PASS | Point 0 | Black | 0.21061 | 0.30529 | 0.22821 | 0.81839
|  |  | Point 1 | White | 176.0229 | 76.7309 | 171.9076 | 71.6525
|  |  | Point 1 | Black | 0.19032 | 0.28007 | 0.21674 | 0.85195
|  |  | Point 2 | White | 170.6234 | 76.118 | 168.843 | 70.7185
|  |  | Point 2 | Black | 0.18665 | 0.27768 | 0.2144 | 0.85107
|  |  | Point 3 | White | 172.9875 | 82.4514 | 161.4881 | 62.1669
|  |  | Point 3 | Black | 0.2006 | 0.30295 | 0.21236 | 0.72674
|  |  | Point 4 | White | 174.7971 | 81.3715 | 164.0273 | 62.1086
|  |  | Point 4 | Black | 0.21402 | 0.305 | 0.22535 | 0.72937

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.077169 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12775 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16721 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.20556 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26326 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34002 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.43721 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.56359 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.67391 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.85691 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1024 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4036 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.7938 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.283 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.8973 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7008 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.7078 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.0095 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.6176 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.5644 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.4655 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 16.3327 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.4159 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 25.5264 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.8423 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 43.225 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 52.7105 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 66.5157 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 85.8662 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 109.5364 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 137.1175 | PASS
31 | 148.75 | 175 | 201.25 | 176.2564 | PASS