## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24090896
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 178.4395 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.13208 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 741.4607 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1716 | 0.4981 | 0.03 | 0.017932 | PASS
Radius Red | 0.4508 | 0.5283 | 0.03 | 0.00036056 | PASS
Radius Green | 0.1231 | 0.5612 | 0.03 | 0.011129 | PASS
Radius Blue | 0.1178 | 0.307 | 0.03 | 0.024557 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.071378 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 63.6262 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 94.1806 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1716 | 0.4981 | 0.017932 | PASS | -0.0054 | 0.0171 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4508 | 0.5283 | 0.00036056 | PASS | -0.0002 | 0.0003 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1231 | 0.5612 | 0.011129 | PASS | 0.0111 | -0.0008 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1178 | 0.307 | 0.024557 | PASS | -0.0052 | 0.024 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 220.3278
Point 1 | 199.0509
Point 2 | 192.9218
Point 3 | 200.5978
Point 4 | 207.3398

## Curves
CURVES
Step | Luminance | Status
1 | 0.10863 | 0
2 | 0.050201 | 0
3 | 0.079445 | 0
4 | 0.064881 | 0
5 | 0.072119 | 0
6 | 0.072119 | 0
7 | 0.10169 | 0
8 | 0.15113 | 0
9 | 0.071594 | 0
10 | 0.11161 | 0
11 | 0.1316 | 0
12 | 0.13163 | 0
13 | 0.17246 | 0
14 | 0.18665 | 0
15 | 0.089369 | 0
16 | 0.13823 | 0
17 | 0.16266 | 0
18 | 0.16263 | 0
19 | 0.21236 | 0
20 | 0.21256 | 0
21 | 0.27135 | 0
22 | 0.27143 | 0
23 | 0.3514 | 0
24 | 0.39898 | 0
25 | 0.19631 | 0
26 | 0.29799 | 0
27 | 0.34878 | 0
28 | 0.34907 | 0
29 | 0.4483 | 0
30 | 0.48449 | 0
31 | 0.23898 | 0
32 | 0.36162 | 0
33 | 0.42262 | 0
34 | 0.42291 | 0
35 | 0.54491 | 0
36 | 0.5452 | 0
37 | 0.69726 | 0
38 | 0.69668 | 0
39 | 0.88697 | 0
40 | 0.8861 | 0
41 | 1.1386 | 0
42 | 1.1377 | 0
43 | 1.4523 | 0
44 | 1.4514 | 0
45 | 1.8501 | 0
46 | 1.8513 | 0
47 | 2.3761 | 0
48 | 2.3577 | 0
49 | 3.0062 | 0
50 | 3.0179 | 0
51 | 3.8205 | 0
52 | 3.8555 | 0
53 | 4.8303 | 0
54 | 4.9033 | 0
55 | 6.2313 | 0
56 | 6.2692 | 0
57 | 7.8599 | 0
58 | 7.8686 | 0
59 | 10.0459 | 0
60 | 9.8942 | 0
61 | 12.5414 | 0
62 | 13.0638 | 0
63 | 16.213 | 0
64 | 15.8657 | 0
65 | 21.0755 | 0
66 | 22.5289 | 0
67 | 11.9109 | 0
68 | 17.4476 | 0
69 | 20.5618 | 0
70 | 21.0755 | 0
71 | 27.0266 | 0
72 | 30.9083 | 0
73 | 15.2615 | 0
74 | 23.1039 | 0
75 | 26.653 | 0
76 | 26.9594 | 0
77 | 33.1849 | 0
78 | 31.6672 | 0
79 | 44.5384 | 0
80 | 46.1436 | 0
81 | 23.673 | 0
82 | 36.1327 | 0
83 | 40.0145 | 0
84 | 38.3509 | 0
85 | 51.3972 | 0
86 | 51.4555 | 0
87 | 65.1147 | 0
88 | 64.9396 | 0
89 | 83.327 | 0
90 | 82.4806 | 0
91 | 105.0417 | 0
92 | 105.1584 | 0
93 | 133.2357 | 0
94 | 133.8486 | 0
95 | 207.4858 | 0
96 | 104.4579 | 0
97 | 155.9719 | 0
98 | 181.4807 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 192.6883
150 | 200 | 178.4395 | PASS | Point 1 | 174.0966
|  |  |  | Point 2 | 168.8139
|  |  |  | Point 3 | 175.264
Uniformity Test |  |  |  | Point 4 | 181.3348
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.13208 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 192.6007 | White
100 | 741.4607 | PASS | 0.25976 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.07714
0.05 | 0.1 | 0.071378 | PASS | Point 1 | 0.069726
|  |  |  | Point 2 | 0.067596
|  |  |  | Point 3 | 0.070281
|  |  |  | Point 4 | 0.072149

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 192.7758 | 82.3055 | 178.1827 | 70.7769
50 | 63.6262 | PASS | Point 1 | 174.1841 | 75.8261 | 168.3177 | 70.952
|  |  | Point 2 | 169.0765 | 74.7754 | 165.3407 | 70.1932
|  |  | Point 3 | 175.1473 | 81.605 | 163.7354 | 63.6262
|  |  | Point 4 | 181.3056 | 82.072 | 169.8646 | 64.1808

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 192.3964 | 82.2471 | 177.9784 | 70.7185
20 | 94.1806 | PASS | Point 0 | Black | 0.26005 | 0.31696 | 0.278 | 0.74279
|  |  | Point 1 | White | 173.8631 | 75.797 | 168.055 | 70.8936
|  |  | Point 1 | Black | 0.24482 | 0.29183 | 0.25885 | 0.72586
|  |  | Point 2 | White | 168.8722 | 74.8046 | 165.3115 | 70.2224
|  |  | Point 2 | Black | 0.23402 | 0.28381 | 0.25313 | 0.72441
|  |  | Point 3 | White | 174.8846 | 81.5175 | 163.5311 | 63.5679
|  |  | Point 3 | Black | 0.20597 | 0.30412 | 0.22856 | 0.67304
|  |  | Point 4 | White | 181.1889 | 82.0136 | 169.6894 | 64.2392
|  |  | Point 4 | Black | 0.24446 | 0.31667 | 0.25827 | 0.68209

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.072411 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.13169 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16257 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21207 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.271 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34819 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.42174 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.54287 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.69347 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.88522 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1368 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.443 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8446 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3562 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9858 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.8117 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8362 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1496 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.8862 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.9584 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.5064 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 16.7267 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.3341 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 27.0353 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 33.5059 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 41.9408 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 50.6383 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 65.5234 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 82.5098 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 104.0493 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 134.374 | PASS
31 | 148.75 | 175 | 201.25 | 181.2181 | PASS