## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24164573
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 175.4567 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.13057 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 684.4501 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1716 | 0.4975 | 0.03 | 0.017361 | PASS
Radius Red | 0.4513 | 0.5286 | 0.03 | 0.00067082 | PASS
Radius Green | 0.1231 | 0.5615 | 0.03 | 0.011111 | PASS
Radius Blue | 0.1186 | 0.3048 | 0.03 | 0.02224 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.073176 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 53.1192 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 60.9536 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1716 | 0.4975 | 0.017361 | PASS | -0.0054 | 0.0165 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4513 | 0.5286 | 0.00067082 | PASS | 0.0003 | 0.0006 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1231 | 0.5615 | 0.011111 | PASS | 0.0111 | -0.0005 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1186 | 0.3048 | 0.02224 | PASS | -0.0044 | 0.0218 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 252.5495
Point 1 | 228.6459
Point 2 | 221.4952
Point 3 | 233.6367
Point 4 | 235.9133

## Curves
CURVES
Step | Luminance | Status
1 | 0.12287 | 0
2 | 0.056476 | 0
3 | 0.089777 | 0
4 | 0.073141 | 0
5 | 0.073112 | 0
6 | 0.1032 | 0
7 | 0.17059 | 0
8 | 0.080321 | 0
9 | 0.12585 | 0
10 | 0.12588 | 0
11 | 0.16517 | 0
12 | 0.16519 | 0
13 | 0.21577 | 0
14 | 0.21577 | 0
15 | 0.27651 | 0
16 | 0.35082 | 0
17 | 0.17013 | 0
18 | 0.26063 | 0
19 | 0.26075 | 0
20 | 0.33652 | 0
21 | 0.33681 | 0
22 | 0.43371 | 0
23 | 0.434 | 0
24 | 0.55863 | 0
25 | 0.55892 | 0
26 | 0.71361 | 0
27 | 0.71477 | 0
28 | 0.9112 | 0
29 | 0.90886 | 0
30 | 1.1692 | 0
31 | 1.168 | 0
32 | 1.4879 | 0
33 | 1.4882 | 0
34 | 1.898 | 0
35 | 2.3758 | 0
36 | 1.187 | 0
37 | 1.7798 | 0
38 | 1.788 | 0
39 | 2.2803 | 0
40 | 2.2646 | 0
41 | 2.8778 | 0
42 | 2.9038 | 0
43 | 3.6629 | 0
44 | 3.6862 | 0
45 | 4.7136 | 0
46 | 4.6844 | 0
47 | 5.9803 | 0
48 | 5.9949 | 0
49 | 7.568 | 0
50 | 7.6556 | 0
51 | 9.7278 | 0
52 | 9.8154 | 0
53 | 12.4159 | 0
54 | 12.5735 | 0
55 | 15.749 | 0
56 | 15.3812 | 0
57 | 19.7562 | 0
58 | 19.5549 | 0
59 | 24.0087 | 0
60 | 32.1634 | 0
61 | 16.607 | 0
62 | 26.0401 | 0
63 | 25.9029 | 0
64 | 33.0389 | 0
65 | 33.7102 | 0
66 | 42.0575 | 0
67 | 42.4078 | 0
68 | 52.2728 | 0
69 | 52.1852 | 0
70 | 66.72 | 0
71 | 66.4865 | 0
72 | 84.6404 | 0
73 | 83.8524 | 0
74 | 108.7483 | 0
75 | 108.4857 | 0
76 | 136.417 | 0
77 | 136.1251 | 0
78 | 233.4908 | 0
79 | 116.6287 | 0
80 | 174.9722 | 0
81 | 82.5098 | 0
82 | 82.539 | 0
83 | 105.0125 | 0
84 | 104.8666 | 0
85 | 133.4692 | 0
86 | 133.761 | 0
87 | 223.5674 | 0
88 | 111.4627 | 0
89 | 167.3253 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 188.8357
150 | 200 | 175.4567 | PASS | Point 1 | 171.1488
|  |  |  | Point 2 | 165.6909
|  |  |  | Point 3 | 174.9138
Uniformity Test |  |  |  | Point 4 | 176.6942
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.13057 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 188.7189 | White
100 | 684.4501 | PASS | 0.27572 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.078832
0.05 | 0.1 | 0.073176 | PASS | Point 1 | 0.071419
|  |  |  | Point 2 | 0.069113
|  |  |  | Point 3 | 0.073199
|  |  |  | Point 4 | 0.073316

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 188.8941 | 67.3037 | 168.2301 | 58.0225
50 | 53.1192 | PASS | Point 1 | 171.1488 | 62.4296 | 159.6785 | 58.2851
|  |  | Point 2 | 165.6909 | 61.0578 | 156.7891 | 57.0885
|  |  | Point 3 | 174.9722 | 67.3329 | 157.3436 | 53.1775
|  |  | Point 4 | 176.7525 | 66.7492 | 160.3498 | 53.1192

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 188.6022 | 67.3037 | 168.2885 | 58.1684
20 | 60.9536 | PASS | Point 0 | Black | 0.27689 | 0.32076 | 0.29002 | 0.92667
|  |  | Point 1 | White | 170.9153 | 62.4296 | 159.3575 | 58.2851
|  |  | Point 1 | Black | 0.25856 | 0.30062 | 0.27321 | 0.90273
|  |  | Point 2 | White | 165.4574 | 61.1162 | 156.7599 | 57.0885
|  |  | Point 2 | Black | 0.24382 | 0.29052 | 0.26834 | 0.93659
|  |  | Point 3 | White | 174.9138 | 67.3621 | 157.3144 | 53.2943
|  |  | Point 3 | Black | 0.26204 | 0.31434 | 0.26274 | 0.79679
|  |  | Point 4 | White | 176.5774 | 66.7492 | 160.2914 | 53.1775
|  |  | Point 4 | Black | 0.26778 | 0.32047 | 0.27899 | 0.85779

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.073637 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12629 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16549 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21595 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26058 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33652 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.43371 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.55863 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.71448 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.91061 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1675 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4818 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.7821 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.2722 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.8743 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.6775 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.6932 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 5.9803 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.5593 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.4914 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.3225 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.8161 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.4626 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 23.9678 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.6088 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 38.1757 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 52.156 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 66.9535 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 84.1151 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 106.7345 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 137.5553 | PASS
31 | 148.75 | 175 | 201.25 | 174.8846 | PASS