## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24332940
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 179.0991 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.17866 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 906.9363 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1706 | 0.4952 | 0.03 | 0.015576 | PASS
Radius Red | 0.4516 | 0.5284 | 0.03 | 0.00072111 | PASS
Radius Green | 0.1203 | 0.5612 | 0.03 | 0.0083385 | PASS
Radius Blue | 0.1195 | 0.3012 | 0.03 | 0.018533 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.07812 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 57.643 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 66.3079 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1706 | 0.4952 | 0.015576 | PASS | -0.0064 | 0.0142 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4516 | 0.5284 | 0.00072111 | PASS | 0.0006 | 0.0004 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1203 | 0.5612 | 0.0083385 | PASS | 0.0083 | -0.0008 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1195 | 0.3012 | 0.018533 | PASS | -0.0035 | 0.0182 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 226.4861
Point 1 | 194.1768
Point 2 | 189.2735
Point 3 | 207.048
Point 4 | 208.77

## Curves
CURVES
Step | Luminance | Status
1 | 0.10781 | 0
2 | 0.049617 | 0
3 | 0.078745 | 0
4 | 0.078657 | 0
5 | 0.11059 | 0
6 | 0.14996 | 0
7 | 0.070777 | 0
8 | 0.11059 | 0
9 | 0.13052 | 0
10 | 0.13055 | 0
11 | 0.17115 | 0
12 | 0.18533 | 0
13 | 0.088405 | 0
14 | 0.13694 | 0
15 | 0.16128 | 0
16 | 0.16125 | 0
17 | 0.2107 | 0
18 | 0.21061 | 0
19 | 0.26986 | 0
20 | 0.26968 | 0
21 | 0.34849 | 0
22 | 0.34878 | 0
23 | 0.44859 | 0
24 | 0.48157 | 0
25 | 0.23635 | 0
26 | 0.35812 | 0
27 | 0.41941 | 0
28 | 0.41882 | 0
29 | 0.54082 | 0
30 | 0.53995 | 0
31 | 0.69055 | 0
32 | 0.69084 | 0
33 | 0.87997 | 0
34 | 0.87851 | 0
35 | 1.1301 | 0
36 | 1.1298 | 0
37 | 1.4374 | 0
38 | 1.436 | 0
39 | 1.8387 | 0
40 | 1.839 | 0
41 | 2.3402 | 0
42 | 2.3372 | 0
43 | 3.0004 | 0
44 | 2.9828 | 0
45 | 3.8263 | 0
46 | 3.7826 | 0
47 | 4.8654 | 0
48 | 4.8683 | 0
49 | 6.1087 | 0
50 | 6.2459 | 0
51 | 7.9504 | 0
52 | 7.8044 | 0
53 | 10.0168 | 0
54 | 10.0226 | 0
55 | 12.7311 | 0
56 | 12.4772 | 0
57 | 15.8686 | 0
58 | 16.3502 | 0
59 | 20.9208 | 0
60 | 20.1765 | 0
61 | 25.6694 | 0
62 | 27.0032 | 0
63 | 30.9959 | 0
64 | 32.1342 | 0
65 | 42.4953 | 0
66 | 44.3924 | 0
67 | 53.6153 | 0
68 | 54.0531 | 0
69 | 68.8798 | 0
70 | 68.1793 | 0
71 | 87.2088 | 0
72 | 87.5299 | 0
73 | 110.0617 | 0
74 | 111.1708 | 0
75 | 141.2619 | 0
76 | 141.5538 | 0
77 | 206.8729 | 0
78 | 103.524 | 0
79 | 155.0379 | 0
80 | 180.6927 | 0
81 | 82.5098 | 0
82 | 82.539 | 0
83 | 105.0125 | 0
84 | 104.8666 | 0
85 | 133.4692 | 0
86 | 133.761 | 0
87 | 223.5674 | 0
88 | 111.4627 | 0
89 | 167.3253 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 197.7083
150 | 200 | 179.0991 | PASS | Point 1 | 169.6894
|  |  |  | Point 2 | 165.2823
|  |  |  | Point 3 | 180.6635
Uniformity Test |  |  |  | Point 4 | 182.152
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.17866 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 197.6792 | White
100 | 906.9363 | PASS | 0.21796 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.086362
0.05 | 0.1 | 0.07812 | PASS | Point 1 | 0.074104
|  |  |  | Point 2 | 0.072119
|  |  |  | Point 3 | 0.078949
|  |  |  | Point 4 | 0.079066

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 197.6792 | 75.0089 | 171.4406 | 63.743
50 | 57.643 | PASS | Point 1 | 169.8937 | 69.6678 | 159.5326 | 64.7353
|  |  | Point 2 | 165.4574 | 68.3836 | 156.5848 | 63.0425
|  |  | Point 3 | 180.6635 | 75.4759 | 160.0579 | 58.2851
|  |  | Point 4 | 182.2396 | 74.075 | 162.3637 | 57.643

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 197.6208 | 75.0673 | 171.1196 | 63.6554
20 | 66.3079 | PASS | Point 0 | Black | 0.21922 | 0.31405 | 0.22643 | 0.89048
|  |  | Point 1 | White | 169.8354 | 69.7554 | 159.1824 | 64.7353
|  |  | Point 1 | Black | 0.30908 | 0.30062 | 0.2646 | 0.97628
|  |  | Point 2 | White | 165.3407 | 68.5296 | 156.1762 | 63.0717
|  |  | Point 2 | Black | 0.19546 | 0.28001 | 0.21837 | 0.91353
|  |  | Point 3 | White | 180.576 | 75.4759 | 159.7077 | 58.2268
|  |  | Point 3 | Black | 0.2252 | 0.31959 | 0.22468 | 0.79825
|  |  | Point 4 | White | 182.1812 | 74.1333 | 162.0134 | 57.643
|  |  | Point 4 | Black | 0.22453 | 0.30908 | 0.22961 | 0.78015

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.07927 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.13096 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16155 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21061 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26965 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34819 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.41882 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.53966 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.68821 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.87938 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1272 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4348 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8279 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3524 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9916 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7942 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8624 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.1233 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.7898 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.9438 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.9471 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 16.4786 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.8682 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.8456 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 34.7318 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 42.5245 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 54.2574 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 68.5004 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 86.917 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 110.9665 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 141.2911 | PASS
31 | 148.75 | 175 | 201.25 | 180.3133 | PASS