## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S24332943
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 179.1108 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.14998 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 906.3056 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1749 | 0.4989 | 0.03 | 0.018023 | PASS
Radius Red | 0.4526 | 0.5287 | 0.03 | 0.0017464 | PASS
Radius Green | 0.1232 | 0.5619 | 0.03 | 0.0112 | PASS
Radius Blue | 0.1184 | 0.3051 | 0.03 | 0.022574 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.076462 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 56.3297 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 79.2437 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1749 | 0.4989 | 0.018023 | PASS | -0.0021 | 0.0179 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4526 | 0.5287 | 0.0017464 | PASS | 0.0016 | 0.0007 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1232 | 0.5619 | 0.0112 | PASS | 0.0112 | -0.0001 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1184 | 0.3051 | 0.022574 | PASS | -0.0046 | 0.0221 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 238.7152
Point 1 | 216.6503
Point 2 | 205.1217
Point 3 | 222.3708
Point 4 | 220.7947

## Curves
CURVES
Step | Luminance | Status
1 | 0.11502 | 0
2 | 0.05201 | 0
3 | 0.083502 | 0
4 | 0.067858 | 0
5 | 0.075593 | 0
6 | 0.075622 | 0
7 | 0.10676 | 0
8 | 0.16044 | 0
9 | 0.074513 | 0
10 | 0.11762 | 0
11 | 0.13928 | 0
12 | 0.12848 | 0
13 | 0.12851 | 0
14 | 0.16893 | 0
15 | 0.16893 | 0
16 | 0.22085 | 0
17 | 0.25917 | 0
18 | 0.12299 | 0
19 | 0.19091 | 0
20 | 0.225 | 0
21 | 0.20772 | 0
22 | 0.20775 | 0
23 | 0.26606 | 0
24 | 0.26606 | 0
25 | 0.3444 | 0
26 | 0.34469 | 0
27 | 0.44363 | 0
28 | 0.44363 | 0
29 | 0.57089 | 0
30 | 0.66311 | 0
31 | 0.32368 | 0
32 | 0.49267 | 0
33 | 0.57672 | 0
34 | 0.53499 | 0
35 | 0.53528 | 0
36 | 0.68471 | 0
37 | 0.68442 | 0
38 | 0.87326 | 0
39 | 0.87413 | 0
40 | 1.119 | 0
41 | 1.1219 | 0
42 | 1.4278 | 0
43 | 1.4223 | 0
44 | 1.8291 | 0
45 | 1.8288 | 0
46 | 2.3276 | 0
47 | 2.3203 | 0
48 | 2.9683 | 0
49 | 2.9683 | 0
50 | 3.7534 | 0
51 | 3.7767 | 0
52 | 4.8245 | 0
53 | 4.8128 | 0
54 | 6.1262 | 0
55 | 6.1321 | 0
56 | 7.7869 | 0
57 | 7.8365 | 0
58 | 9.8008 | 0
59 | 9.8154 | 0
60 | 12.9208 | 0
61 | 12.3517 | 0
62 | 16.1663 | 0
63 | 16.1838 | 0
64 | 20.7544 | 0
65 | 20.0306 | 0
66 | 25.4709 | 0
67 | 25.8182 | 0
68 | 32.9806 | 0
69 | 34.5275 | 0
70 | 42.2618 | 0
71 | 42.2618 | 0
72 | 53.557 | 0
73 | 53.6445 | 0
74 | 67.7999 | 0
75 | 68.2961 | 0
76 | 86.2457 | 0
77 | 85.8371 | 0
78 | 110.2952 | 0
79 | 111.3459 | 0
80 | 140.9701 | 0
81 | 140.8533 | 0
82 | 221.1742 | 0
83 | 109.7407 | 0
84 | 165.1656 | 0
85 | 192.8926 | 0
86 | 179.0583 | 0
87 | 223.5674 | 0
88 | 111.4627 | 0
89 | 167.3253 | 0
90 | 154.0456 | 0
91 | 179.8171 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 193.5055
150 | 200 | 179.1108 | PASS | Point 1 | 176.1688
|  |  |  | Point 2 | 166.5081
|  |  |  | Point 3 | 180.5176
Uniformity Test |  |  |  | Point 4 | 178.854
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.14998 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 193.3888 | White
100 | 906.3056 | PASS | 0.21338 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.082685
0.05 | 0.1 | 0.076462 | PASS | Point 1 | 0.075242
|  |  |  | Point 2 | 0.071186
|  |  |  | Point 3 | 0.077315
|  |  |  | Point 4 | 0.075885

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 193.3596 | 70.8936 | 172.0244 | 61.7291
50 | 56.3297 | PASS | Point 1 | 176.1688 | 67.187 | 165.1364 | 62.7507
|  |  | Point 2 | 166.6249 | 64.9688 | 158.8321 | 61.1454
|  |  | Point 3 | 180.6343 | 71.1855 | 162.2761 | 56.7966
|  |  | Point 4 | 178.854 | 70.2224 | 162.9182 | 56.3297

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 193.0969 | 70.8061 | 171.82 | 61.6708
20 | 79.2437 | PASS | Point 0 | Black | 0.21417 | 0.27701 | 0.22526 | 0.77519
|  |  | Point 1 | White | 176.0229 | 67.2162 | 165.1364 | 62.7798
|  |  | Point 1 | Black | 0.18878 | 0.26262 | 0.21484 | 0.77461
|  |  | Point 2 | White | 166.5665 | 65.0272 | 158.5986 | 61.1454
|  |  | Point 2 | Black | 0.18589 | 0.25214 | 0.20935 | 0.77081
|  |  | Point 3 | White | 180.3133 | 71.1855 | 162.1594 | 56.7966
|  |  | Point 3 | Black | 0.23405 | 0.27937 | 0.22984 | 0.69201
|  |  | Point 4 | White | 178.5037 | 70.164 | 162.6555 | 56.2713
|  |  | Point 4 | Black | 0.1968 | 0.27359 | 0.21294 | 0.7101

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.076118 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12877 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16896 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.2074 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.2656 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.34411 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.44276 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.53353 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.68238 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.87034 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1167 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.424 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.8145 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.3308 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.942 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.7884 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.8333 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 6.17 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.8949 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 10.0343 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.3604 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.5709 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.1123 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.6296 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 32.3093 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 40.4523 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 53.5861 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 67.6831 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 85.8079 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 109.945 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 140.0653 | PASS
31 | 148.75 | 175 | 201.25 | 178.5621 | PASS