## ATP
5.6.2 Display Test | Min |  | Max | Result | [Result] | Serial Number
S25001608
5.6.2.1 Luminance Test
L_avg | 150 |  | 200 | 176.1922 | PASS
5.6.2.2 Uniformity Test
UNIFORMITY |  |  | 0.65 | 0.15385 | PASS
5.6.2.3 Contrast on-axis
Min Dark Ambient Contrast | 100 |  |  | 943.2321 | PASS
5.6.2.4 Chromaticity Coordinates (DAY) | u' | v' | R
Radius White | 0.1726 | 0.4984 | 0.03 | 0.017948 | PASS
Radius Red | 0.4518 | 0.5286 | 0.03 | 0.001 | PASS
Radius Green | 0.1221 | 0.5619 | 0.03 | 0.0101 | PASS
Radius Blue | 0.1185 | 0.3033 | 0.03 | 0.020793 | PASS
5.6.2.5 Luminance Dimming Test
L_avg | 0.05 |  | 0.1 | 0.075359 | PASS
5.6.2.6 Luminance Viewing Angle
L_min | 50 |  |  | 55.6584 | PASS
5.6.2.7 Contrast for Secondary Viewing angle
Min Dark ambient contrast | 20 |  |  | 66.0046 | PASS
*Curves
Step Encoder 0 |  |  |  |  | PASS
Step Encoder 1 |  |  |  |  | PASS
Step Encoder 2 |  |  |  |  | PASS
Step Encoder 3 |  |  |  |  | PASS
Step Encoder 4 |  |  |  |  | PASS
Step Encoder 5 |  |  |  |  | PASS
Step Encoder 6 |  |  |  |  | PASS
Step Encoder 7 |  |  |  |  | PASS
Step Encoder 8 |  |  |  |  | PASS
Step Encoder 9 |  |  |  |  | PASS
Step Encoder 10 |  |  |  |  | PASS
Step Encoder 11 |  |  |  |  | PASS
Step Encoder 12 |  |  |  |  | PASS
Step Encoder 13 |  |  |  |  | PASS
Step Encoder 14 |  |  |  |  | PASS
Step Encoder 15 |  |  |  |  | PASS
Step Encoder 16 |  |  |  |  | PASS
Step Encoder 17 |  |  |  |  | PASS
Step Encoder 18 |  |  |  |  | PASS
Step Encoder 19 |  |  |  |  | PASS
Step Encoder 20 |  |  |  |  | PASS
Step Encoder 21 |  |  |  |  | PASS
Step Encoder 22 |  |  |  |  | PASS
Step Encoder 23 |  |  |  |  | PASS
Step Encoder 24 |  |  |  |  | PASS
Step Encoder 25 |  |  |  |  | PASS
Step Encoder 26 |  |  |  |  | PASS
Step Encoder 27 |  |  |  |  | PASS
Step Encoder 28 |  |  |  |  | PASS
Step Encoder 29 |  |  |  |  | PASS
Step Encoder 30 |  |  |  |  | PASS
Step Encoder 31 |  |  |  |  | PASS

## ChromaticityCoordinates
DAY CHROMATICITY COORDINATES | [10 empty] | PASS/FAIL
| REQUIRED |  |  | RESULT
DAY | u' | v' | r | u' | v' | r
white | 0.177 | 0.481 | 0.03 | 0.1726 | 0.4984 | 0.017948 | PASS | -0.0044 | 0.0174 | #VALUE! | PASS
red | 0.451 | 0.528 | 0.03 | 0.4518 | 0.5286 | 0.001 | PASS | 0.0008 | 0.0006 | #VALUE! | PASS
green | 0.112 | 0.562 | 0.03 | 0.1221 | 0.5619 | 0.0101 | PASS | 0.0101 | -0.0001 | #VALUE! | PASS
blue | 0.123 | 0.283 | 0.03 | 0.1185 | 0.3033 | 0.020793 | PASS | -0.0045 | 0.0203 | #VALUE! | PASS

## CalibPosition
Luminance
Luminance
Point 0 | 237.431
Point 1 | 215.045
Point 2 | 203.4872
Point 3 | 214.4321
Point 4 | 212.8561

## Curves
CURVES
Step | Luminance | Status
1 | 0.11234 | 0
2 | 0.051514 | 0
3 | 0.081926 | 0
4 | 0.066749 | 0
5 | 0.074279 | 0
6 | 0.074279 | 0
7 | 0.1046 | 0
8 | 0.15632 | 0
9 | 0.07352 | 0
10 | 0.11511 | 0
11 | 0.13589 | 0
12 | 0.12565 | 0
13 | 0.12565 | 0
14 | 0.16479 | 0
15 | 0.16484 | 0
16 | 0.21534 | 0
17 | 0.21534 | 0
18 | 0.2759 | 0
19 | 0.32222 | 0
20 | 0.15621 | 0
21 | 0.23944 | 0
22 | 0.28127 | 0
23 | 0.25999 | 0
24 | 0.26008 | 0
25 | 0.33593 | 0
26 | 0.33593 | 0
27 | 0.43254 | 0
28 | 0.43254 | 0
29 | 0.55717 | 0
30 | 0.55804 | 0
31 | 0.71331 | 0
32 | 0.71361 | 0
33 | 0.90886 | 0
34 | 0.90886 | 0
35 | 1.1672 | 0
36 | 1.1637 | 0
37 | 1.4815 | 0
38 | 1.4862 | 0
39 | 1.8948 | 0
40 | 2.1901 | 0
41 | 1.0898 | 0
42 | 1.6438 | 0
43 | 1.9155 | 0
44 | 1.7745 | 0
45 | 1.7725 | 0
46 | 2.2762 | 0
47 | 2.2675 | 0
48 | 2.8702 | 0
49 | 2.895 | 0
50 | 3.657 | 0
51 | 3.657 | 0
52 | 4.6757 | 0
53 | 4.699 | 0
54 | 5.9277 | 0
55 | 5.9277 | 0
56 | 7.5739 | 0
57 | 7.6439 | 0
58 | 9.6315 | 0
59 | 9.7979 | 0
60 | 12.3166 | 0
61 | 12.0481 | 0
62 | 15.7781 | 0
63 | 15.2353 | 0
64 | 19.6278 | 0
65 | 20.2261 | 0
66 | 25.3483 | 0
67 | 26.0488 | 0
68 | 31.6088 | 0
69 | 31.171 | 0
70 | 38.1466 | 0
71 | 47.5446 | 0
72 | 22.8733 | 0
73 | 36.804 | 0
74 | 43.5752 | 0
75 | 39.6643 | 0
76 | 36.1327 | 0
77 | 49.3833 | 0
78 | 61.2622 | 0
79 | 29.7701 | 0
80 | 45.268 | 0
81 | 53.2943 | 0
82 | 52.4771 | 0
83 | 67.2745 | 0
84 | 67.3621 | 0
85 | 86.1581 | 0
86 | 85.049 | 0
87 | 108.4565 | 0
88 | 109.0694 | 0
89 | 138.8978 | 0
90 | 137.9639 | 0
91 | 215.8039 | 0
92 | 107.8144 | 0
93 | 161.6632 | 0
94 | 188.5146 | 0
95 | 175.0597 | 0

## LuminanceAndUniformity
Luminance Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 193.0677
150 | 200 | 176.1922 | PASS | Point 1 | 175.0889
|  |  |  | Point 2 | 165.4866
|  |  |  | Point 3 | 174.2425
Uniformity Test |  |  |  | Point 4 | 173.0751
| REQUIRED UNIFORMITY MAX | MEASURED | RESULT
| 0.65 | 0.15385 | PASS

## ContrastOnAxis
Contrast Test |  |  | Measured Values (fL)
REQUIRED CONTRAST MIN | MEASURED | RESULT | 193.0093 | White
100 | 943.2321 | PASS | 0.20463 | Black

## LuminanceDimming
Luminance Dimming Test |  |  |  |  | Measured Values (fL)
REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE MAX | MEASURED | RESULT | Point 0 | 0.082772
0.05 | 0.1 | 0.075359 | PASS | Point 1 | 0.074805
|  |  |  | Point 2 | 0.070806
|  |  |  | Point 3 | 0.074717
|  |  |  | Point 4 | 0.073696

## LuminanceViewingAngle
Luminance Viewing Angle |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED LUMINANCE MIN | MEASURED MIN | RESULT | Point 0 | 193.2428 | 67.508 | 171.6449 | 59.4234
50 | 55.6584 | PASS | Point 1 | 175.2932 | 61.1162 | 162.0426 | 58.8105
|  |  | Point 2 | 165.6909 | 60.3866 | 157.8982 | 58.1392
|  |  | Point 3 | 174.3884 | 68.6171 | 157.3144 | 57.0885
|  |  | Point 4 | 173.221 | 66.4865 | 160.1163 | 55.6584

## ContrastForSecondaryViewing
Contrast Secondary Viewing Angle |  |  |  |  | Measured Values (fL) (0,0) | Measured Values (fL) (0,55) | Measured Values (fL) (15,0) | Measured Values (fL) (15,55)
REQUIRED CONTRAST MIN | MEASURED MIN | RESULT | Point 0 | White | 193.0385 | 67.5372 | 171.7033 | 59.4526
20 | 66.0046 | PASS | Point 0 | Black | 0.20556 | 0.29858 | 0.22281 | 0.868
|  |  | Point 1 | White | 175.0889 | 61.1162 | 162.0134 | 58.8397
|  |  | Point 1 | Black | 0.19024 | 0.27263 | 0.21151 | 0.86975
|  |  | Point 2 | White | 165.3699 | 60.4157 | 157.8398 | 58.1976
|  |  | Point 2 | Black | 0.18939 | 0.26968 | 0.21254 | 0.88172
|  |  | Point 3 | White | 174.2133 | 68.6755 | 157.2852 | 57.1761
|  |  | Point 3 | Black | 0.1879 | 0.30091 | 0.20533 | 0.84407
|  |  | Point 4 | White | 172.9583 | 66.4865 | 160.0871 | 55.6876
|  |  | Point 4 | Black | 0.19905 | 0.29478 | 0.21615 | 0.80817

## TestCurves
Test Curves
ENCODER POSITION | REQUIRED LUMINANCE MIN | REQUIRED LUMINANCE | REQUIRED LUMINANCE MAX | MEASURED | RESULT
0 | 0.06375 | 0.075 | 0.08625 | 0.075096 | PASS
1 | 0.10843 | 0.12756 | 0.14669 | 0.12649 | PASS
2 | 0.13795 | 0.16229 | 0.18663 | 0.16572 | PASS
3 | 0.1755 | 0.20648 | 0.23745 | 0.21627 | PASS
4 | 0.22329 | 0.26269 | 0.3021 | 0.26075 | PASS
5 | 0.28408 | 0.33421 | 0.38435 | 0.33681 | PASS
6 | 0.36143 | 0.42521 | 0.48899 | 0.434 | PASS
7 | 0.45983 | 0.54098 | 0.62212 | 0.55833 | PASS
8 | 0.58503 | 0.68827 | 0.79151 | 0.7139 | PASS
9 | 0.74431 | 0.87566 | 1.007 | 0.90828 | PASS
10 | 0.94696 | 1.1141 | 1.2812 | 1.1645 | PASS
11 | 1.2048 | 1.4174 | 1.63 | 1.4841 | PASS
12 | 1.5328 | 1.8033 | 2.0738 | 1.7807 | PASS
13 | 1.9501 | 2.2943 | 2.6384 | 2.2602 | PASS
14 | 2.4811 | 2.9189 | 3.3568 | 2.9014 | PASS
15 | 3.1566 | 3.7136 | 4.2707 | 3.6833 | PASS
16 | 4.016 | 4.7247 | 5.4334 | 4.7194 | PASS
17 | 5.1094 | 6.0111 | 6.9128 | 5.954 | PASS
18 | 6.5005 | 7.6477 | 8.7949 | 7.7081 | PASS
19 | 8.2704 | 9.7299 | 11.1894 | 9.7453 | PASS
20 | 10.5222 | 12.379 | 14.2359 | 12.0569 | PASS
21 | 13.387 | 15.7494 | 18.1118 | 15.7023 | PASS
22 | 17.0318 | 20.0374 | 23.043 | 20.2612 | PASS
23 | 21.6689 | 25.4929 | 29.3168 | 26.0634 | PASS
24 | 27.5686 | 32.4337 | 37.2987 | 31.7256 | PASS
25 | 35.0746 | 41.2642 | 47.4538 | 36.2203 | PASS
26 | 44.6241 | 52.499 | 60.3738 | 52.8857 | PASS
27 | 56.7737 | 66.7926 | 76.8115 | 67.187 | PASS
28 | 72.2312 | 84.9778 | 97.7245 | 85.2241 | PASS
29 | 91.8972 | 108.1143 | 124.3314 | 108.544 | PASS
30 | 116.9175 | 137.55 | 158.1825 | 138.606 | PASS
31 | 148.75 | 175 | 201.25 | 175.264 | PASS